Understanding Linear Power

Supply Operation

Application Note 1554

Contents

Overview of linear power supply operation / 2

Output characteristics / 3

Extending the voltage and current range / 5

Remote programming / 6

This application note describes the basic operation

of linear power supplies. You will also find

information to help you better understand output

characteristics of a linear power supply compared

to an ideal supply.

Overview of linear power supply operation

The power supply shown in

Figure 1

has two ranges,

allowing more voltage at a lower current or more

The basic design model for a power supply consists

current at a lower voltage. Single-range supplies can

of a control element in series with a rectifier and

output maximum power only at full-scale voltages

load device.

Figure 1

shows a simplified schematic

and full-scale current. A linear supply can provide

of a series-regulated supply with the phase-con-

output power that is close to maximum at full scale

trolled pre-regulator depicted as a power switch and

for both ranges. The pre-regulator in this power

the series element depicted as a variable resistor.

supply uses solid-state transformer tap switches

The phase-controlled pre-regulator minimizes the

on the secondary winding of the power transformer.

power dissipated at the series element by maintain-

This technique is very effective in reducing the

ing a low and constant voltage drop across the series

power dissipated in the series element.

element. Feedback control circuits continuously

In terms of performance, a linear regulated supply

monitor the output and adjust the series resistance

has very precise regulating properties and responds

to maintain a constant output voltage. The variable

quickly to variations of the line and load. Hence,

resistance series element of the supply shown

its line and load regulation and transient recovery

in

Figure 1

is actually produced by one or more

time are superior to supplies using other regulation

power transistor operating in the linear (class A)

techniques. A linear power supply also exhibits

mode; supplies with this type of regulator are often

low ripple and noise, tolerates ambient temperature

called linear power supplies. Linear power supplies

changes, and is highly reliable due to its circuit

have many advantages. Because they provide suffi-

simplicity.

cient power with stable regulation and little noise,

they usually are the simplest, most effective

The linear regulator is controlled by a DAC driven by

solution for providing bench power.

a digital circuit that provides a voltage proportional to

the program voltage. The power supply sends back to

Phase

control

the control circuits a voltage representing the output

Rectifier

R an g e  1

Series

at the terminals. The control circuits receive informa-

element

tion from the front panel and send information to the

AC Input

Ra n g e  2

Pre-regulator

display. Similarly, the control circuits “talks” to the

Filter

remote interface for input and output with the GPIB,

Feedback

DC

Power

P

output

control

transformer

L

RS-232, USB,  or LAN  interfaces. The remote inter-

I

L

face is at earth ground and is optically isolated from

the control circuit and the power supply.

Figure 1.

Diagram of simple series power supply

+

GPIB

Power

Optical

Control

supply

isolation

display

output

RS-232

Figure 2.

Block diagram of the power supply showing the optical isolation

2

Output characteristics

The output of this power supply can operate in

either constant-voltage (CV) mode or constant-

An ideal constant-voltage power supply would have

current (CC) mode. Under certain fault conditions,

zero output impedance at all frequencies. Thus,

the power supply cannot operate in either CV or

as shown in

Figure 3

, the voltage would remain

CC mode and becomes unregulated.

perfectly constant in spite of any changes in output

Figure 5

shows the operating modes of the output

current demanded by the load.

of this power supply. The operating point of

one supply will be either above or below the line

(V)

R

= R

. This line represents a load where the

L

C

E

output voltage and the output current are equal to

S

the voltage and current setting. When the load R

L

is greater than R

, the output voltage will domi-

C

E

~

nate since the current will be less then the current

E

0

Ou t

Ou t

setting. The power supply is said to be in constant

voltage mode. The load at point 1 has a relatively

high resistance value (compared to R

), the

C

output voltage is at the voltage setting, and the

output current is less than the current setting.

I

(A)

Ou t

In this case, the power supply is in the constant

00Figure 3.

Ideal constant-voltage power supply

voltage mode and the current setting acts as a

current limit.

(V)

V

O ut

~

CV operating line

I

0

O u t

V

M ax

V

= Voltage setting value

S

1

I

=Current setting value

E

V

S

S

O u t

R

=V

/I

R

=

R

C

S

S

I

L

C

R

=Load resistance

L

S

R

>

R

L

C

CC

2

operating

line

I

(A)

R

<

R

Ou t

L

C

Figure 4.

Ideal constant-current power supply

I

I

I

O ut

S

Max

Figure 5.

Output characteristics

The ideal constant-current power supply exhibits

infinite output impedance at all frequencies. Thus,

as

Figure 4

indicates, the ideal constant-current

power supply would accommodate a load resistance

change by altering its output voltage by just the

amount necessary to maintain its output current

at a constant value.

3

When the load RL is less than RC, the output

Normal mode voltage noise is in the form of ripple

current will dominate since the voltage will be

related to the line frequency plus some random

less than the set voltage. The power supply is said

noise. Both of these are very low in a quality bench

to be in constant current mode. The load at point

power supply. Careful lead layout and keeping the

2 has a relatively low resistance, the output voltage

power supply circuitry away from power devices and

is less than the voltage setting, and the output

other noise sources will keep these values low.

current is at the current setting. The supply is in

Common mode noise can be a problem for very sensi-

constant current mode and the voltage setting

tive circuitry that is referenced to earth ground. When

acts as a voltage limit.

a circuit is referenced to earth ground, a low-level cur-

rent related to the AC line will flow from the output

Unregulated state

terminals to earth ground. Any impedance to earth

If the power supply should go into a mode of opera-

ground will create a voltage drop equal to the current

tion that is neither CV nor CC, the power supply

flow multiplied by the impedance. To minimize this

is unregulated. In this mode the output is not

effect, the output terminal can be grounded at the

predictable. The unregulated condition may be

output terminal. Alternately, any impedance to earth

the result of the AC line voltage below the specifi-

ground should have complementary impedance to

cations. The unregulated condition may occur

earth ground to cancel any generated voltages. If the

momentarily. For example, when the output is

circuit is not referenced to earth ground, common

programmed for a large voltage step, the output

mode power line noise is typically not a problem.

capacitor or a large capacitive load will charge up

The output will also change due to changes in the

at the current limit setting. During the ramp up of

load. As the load increases, the output current will

the output voltage the power supply will be in

cause a small drop in the output voltage of the power

the unregulated mode. During the transition from

supply due to the output impedance (R). Any resist-

CV to CC, as when the output is shorted, the unreg-

ance in the connecting wire will add to this resistance

ulated state may occur briefly during the transition.

and increase the voltage drop. You can minimize the

voltage drop by using the largest possible hook up

Unwanted signals

wire. Using the remote sense leads at the load will

An ideal power supply has a perfect DC output with

compensate for lead resistance in the load leads.

no signals across the terminals or from the terminals

to earth ground. An actual power supply has finite

Normal

mode

noise across the output terminals and a finite

noise

R

current will flow through any impedance connected

+

from either terminal to earth ground. The first is

<5mV

*

pp

called normal mode voltage noise and the second

Output

<0.5mV

*

Output

rms

voltage

terminal

common mode current noise.

Figure 6

shows

<8mV

**

pp

<1mV

**

a simplified diagram of common mode and normal

rms

mode sources of noise.

*

For the E3646A and the E3648A power supply

Common

mode

**

For the E3647A and the E3649A power supply

noise

<1.5

µ

A

rms

Figure 6.

Simplified diagram of common mode and normal mode

4

When the load changes very rapidly, as when a relay

When you use series connection, the output voltage

contact is closed, the inductance in the hook-up wire

is the sum of the voltages of the individual power

and in the power supply output will cause a spike to

supplies. The current is the current of any one

appear at the load. The spike is a function of the rate

power supply. You must adjust each of the individual

of change of the load current. When you expect very

power supplies to obtain the total output voltage.

rapid changes in load, you can minimize these volt-

Parallel connections

age spikes by using a capacitor with a low series

resistance, in parallel with the power supply and

If you have power supplies capable of CV/CC auto-

close to the load.

matic crossover operation, you can connect two

or more of them in parallel to obtain a total output

Extending the voltage and current range

current greater than that available from one power

A power supply may be able to provide voltages and

supply. The total output current is the sum of the

current greater than its rated maximum outputs if

output currents of the individual power supplies.

the power-line voltage is at or above its nominal

You can set the output of each power supply

value. Operation can be extended typically up to 3%

separately. The output voltage controls of one power

over the rated output without damage to the power

supply should be set to the desired output voltage;

supply, but performance cannot be guaranteed to

the other power supply should be set for a slightly

meet specifications in this region. If the power-line

higher output voltage. The supply with the higher

voltage is maintained in the upper end of the input

output voltage setting will deliver its constant

voltage range, the power supply will probably operate

current output and drop its output voltage until it

within its specifications. A power supply is more

equals the output of the other supply, and the other

likely to stay within specifications if only one of

supply will remain in constant voltage operation and

the voltage or current outputs is exceeded.

only deliver that fraction of its rated output current

that is necessary to fulfill the total load demand.

Series connections

If you need a higher voltage than you can get from

a single supply, you can operate two or more power

supplies in series up to the output isolation rating

of any one supply. You can operate series-connected

power supplies with one load across both power

supplies or with a separate load for each power

supply. A quality bench power supply has a reverse

polarity diode connected across the output terminals

so that if you operate it in series with other power

supplies, damage will not occur if the load is short-

circuited or if one power supply is turned on

separately from its series partners.

5

Remote programming

When this exponential rise reaches the newly

programmed voltage level, the constant voltage

During remote programming a constant-voltage

amplifier resumes its normal regulating action

regulated power supply is called upon to change

and holds the output constant. Thus, the rise

its output voltage rapidly. The most important

time can be determined approximately using the

factor limiting the speed of output voltage change

formula shown in

Figure 7

.

is the output capacitor and load resistor.

If no load resistor is attached to the power supply

output terminal, then the output voltage will rise

(V )

Asymptotic level I

R

L

L

linearly at a rate of C

/ I

when programmed

O

L

E

= New E

upward, and T

= C

(E

– E

)/ I

, the shortest

2

Ou t

R

O

2

1

L

I

TC = R

C

possible up-programming time.

+

L

C

E

E

L

O

R

O

O ut

L

O u t

(

E

= Initial E

1

O u t

I

R

--E

T

T

= Rise time = R

C

LOG

L

L

1

( V)

e

L

O

R

R

I

R

--E

L

L

2

0

Time

Equivalent circuit for

( S)

E

= Initial E

t

< t < t

I

1

O u t

1

2

t

,  new E

is reached

+

L

O u t

TC = R

C

C

E

2

R

E

L

O

t

,  new E

is programmed

O

O u t

L

Ou t

O u t

E

= New E

1

(

2

O u t

E

T

= Fall time = R

C

LOG

1

T

e

L

O

F

Figure 7.

Speed of response — programming up (full load)

E

F

2

( S)

0

Equivalent circuit for

Asymptotic level = 0

Time

t

< t < t

1

2

t

,  new E

is reached

In

Figure 7

, you can see the equivalent circuit and

O u t

2

t

,  new E

is programmed

Ou t

the nature of the output voltage waveform when

1

the supply is being programmed upward. When

Figure 8.

Speed of response — programming down

the new output is programmed, the power supply

regulator circuit senses that the output is less

Figure 8

shows that when the power supply is

than desired and turns on the series regulator

programmed down, the regulator senses that the

to its maximum value I

, the current limit or

output voltage is higher than desired and turns

L

constant current setting.

off the series transistors entirely. Since the control

circuits cannot cause the series regulator transis-

This constant current I

charges the output capacitor

L

tors to conduct backwards, the output capacitor

C

and parallel load resistor R

. The output there-

O

L

can only be discharged through the load resistor

fore rises exponentially with a time constant R

C

L

L

and internal current source (I

).

S

towards voltage level I

R

, a value higher than the

L

L

new output voltage being programmed.

The output voltage decays linearly with slope of

I

/C

with no load and stops falling when it

S

O

reaches the new output voltage, which has been

demanded. If full load is connected, the output

voltage will fall exponentially faster.

Since up-programming speed is aided by the con-

duction of the series-regulating transistor, while

down programming normally has no active element

aiding in the discharge of the output capacitor,

bench power supplies normally program upward

more rapidly than downward.

6

Glossary

Linear supply

Auto range

– To regulate the output, one or more power

– When the power supply’s regulator uses a variable

voltage source, the lower the voltage source the more current can

transistors are used to reduce the output voltage to the proper

be output by the power supply.  The power supply picks the lowest

value. The power transistors operate in linear (class A) mode.

internal voltage source to meet the output voltage requirement in

Normal mode noise

– Noise the flows in the same direction

order to supply the maximum current. An auto-range supply can

as the power from the supply

cover more current combinations than a single- or dual-range supply

Output voltage

– Voltage measured at the output terminals. Using

Common mode noise

– Noise that travels through a ground loop

remote sense will increase the output voltage above the program

Constant current mode

voltage to compensate for the voltage drop in the leads.

– A CV/CC power supply cannot be

set to output a constant current.  Instead, the current limit needs

OVP

– Over voltage protection, an independent circuit used

to be set below the current drawn by the load.  Once a power

to clamp the output voltage, protecting the load

supply has switched from regulating the output voltage to limiting

Program voltage

the current, it is in a constant current mode.

– Voltage set via the front panel, remote

interface or analog programming terminals

Constant voltage mode

– The power supply will output the

Ripple and noise

–The AC component of a power supply output;

programmed voltage with a varying load unless the output current

rises above the programmed current setting.

both the RMS value and the peak value should be specified

Down programming

Settling time

– Setting the power supply's output voltage

–Time required for the output to reach the

programmed voltage

to a smaller value. The power supply is forced to sink current.

Single range

Dual range

– The output can be characterized as having a

– Output voltage is generated from two internal

voltage sources. The lower-voltage source can generate more

max voltage and a max current. A single range supply can output

any voltage/current combination within these limits.

current. A dual power supply can cover more voltage current

combinations than a single-range supply.

Speed of response

– The time it takes to change from one

Ideal constant-current power supply

– A supply with infinite

voltage to a new program voltage

output impendence; the supply will output the program current

Transient response

– The time required for the output voltage

under any load

to recover from a change in load

Ideal constant voltage power supply

– A supply with

Unregulated state

– If the power supply should go into a mode

zero output impendence; the supply will output the program

of operation that is neither CV nor CC, the power supply is

voltage under any load

unregulated. In this mode the output is not predictable.

Related Agilent literature

Publication title Publication  Publication  Web address

type number

E3631A, E3632A, E3633A,

and E3634A Bench Power  Data sheet 5968-9726EN http://cp.literature.agilent.com/litweb/pdf/5968-9726EN.pdf

Supplies

E3640A - E3649A  Data sheet 5968-7355EN http://cp.literature.agilent.com/litweb/pdf/5968-7355EN.pdf

DC Power Supplies

E3610A - E3630A  Data sheet 5968-9727EN http://cp.literature.agilent.com/litweb/pdf/5968-9727EN.pdf

Manual Power Supplies

7

www.agilent.com

By internet, phone, or fax, get assistance with

all your test & measurement needs.

Online assistance:

www.agilent.com/find/emailupdates

www.agilent.com/find/assist

Get the latest information on the products and applications you select.

Phone or Fax

Agilent T&M Software and Connectivity

United States:

Agilent’s Test and Measurement software and connectivity products, solutions and developer

(tel) 800 829 4444

network allows you to take time out of connecting your instruments to your computer with tools

(fax) 800 829 4433

based on PC standards, so you can focus on your tasks, not on your connections.

Canada:

www.agilent.com/find/connectivity

Visit

for more information.

(tel)  877 894 4414

(fax) 800 746 4866

China:

(tel) 800 810 0189

(fax) 800 820 2816

Europe:

(tel)  (31 20) 547 2111

(fax) (31 20) 547 2390

Japan:

(tel)  (81) 426 56 7832

(fax) (81) 426 56 7840

Korea:

(tel)  (82 2) 2004 5004

(fax) (82 2) 2004 5115

Latin America:

(tel) (650) 752 5000

Taiwan

:

(tel) 0800 047 866

(fax) 0800 286 331

Other Asia Pacific Countries:

(tel) (65) 6375 8100

(fax) (65) 6836 0252

Email: tm_ap@agilent.com

Product specifications and descriptions in this

document subject to change without notice.

02/04/2005

© Agilent Technologies, Inc. 2005

Printed in USA February 4, 2005

5989-2291EN